## measurement graphs
Correction filter parameters
[9 empty] | 90 deg offset |  | freq | 9300 | 5500
[9 empty] | -10 |  | Q | 0.707 | 1.2
[11 empty] | gain | -7.5 | 5.5
Frequency_Response | FrequencyResponse | [10 empty] | Correction filter responses
Hz | B&K 4191 | OpenRefMic prototype | Dayton EMM-6 | B&K 4191 90deg off axis | OpenRefMic 90 deg off axis | OpenRefMic relative-response-equalized | OpenRefMic relative-response-raw | Dayton EMM-6 relative-response | B&K 4191 90deg off axis (offset) | OpenRefMic 90 deg off axis (offset) | response-curve_modeled | Shelf | Peak | OpenRefMic modeled response | OpenRefMic correction curve
10 | 57.867 | 57.7194 | 58.3521 | 58.0144 | 58.0024 | -0.14761 | -0.1476 | 0.4851 | -9.8526 | -9.8646 |  | -1.9543e-09 | 1.2129e-05 | 1.2127e-05 | -1.2127e-05
10.2932 | 58.0584 | 57.904 | 58.5926 | 58.203 | 58.1997 | -0.15441 | -0.1544 | 0.5342 | -9.8554 | -9.8587 |  | -2.071e-09 | 1.285e-05 | 1.2848e-05 | -1.2848e-05
10.595 | 58.2566 | 58.0981 | 58.8368 | 58.3983 | 58.4038 | -0.15851 | -0.1585 | 0.5802 | -9.8583 | -9.8528 |  | -2.1946e-09 | 1.3615e-05 | 1.3613e-05 | -1.3613e-05
10.9057 | 58.4629 | 58.3038 | 59.0853 | 58.6015 | 58.6161 | -0.15911 | -0.1591 | 0.6224 | -9.8614 | -9.8468 |  | -2.3257e-09 | 1.4425e-05 | 1.4423e-05 | -1.4423e-05
11.2254 | 58.6767 | 58.52 | 59.3374 | 58.8119 | 58.8357 | -0.15672 | -0.1567 | 0.6607 | -9.8648 | -9.841 |  | -2.4646e-09 | 1.5283e-05 | 1.5281e-05 | -1.5281e-05
11.5546 | 58.8983 | 58.7476 | 59.593 | 59.03 | 59.0632 | -0.15072 | -0.1507 | 0.6947 | -9.8683 | -9.8351 |  | -2.6119e-09 | 1.6193e-05 | 1.619e-05 | -1.619e-05
11.8934 | 59.1282 | 58.9873 | 59.8522 | 59.2562 | 59.299 | -0.14092 | -0.1409 | 0.724 | -9.872 | -9.8292 |  | -2.7679e-09 | 1.7157e-05 | 1.7154e-05 | -1.7154e-05
12.2421 | 59.3662 | 59.2387 | 60.1146 | 59.4903 | 59.5427 | -0.12752 | -0.1275 | 0.7484 | -9.8759 | -9.8235 |  | -2.9334e-09 | 1.8177e-05 | 1.8174e-05 | -1.8174e-05
12.601 | 59.6123 | 59.5021 | 60.3799 | 59.7324 | 59.7946 | -0.11022 | -0.1102 | 0.7676 | -9.8799 | -9.8177 |  | -3.1088e-09 | 1.9259e-05 | 1.9256e-05 | -1.9256e-05
12.9705 | 59.8668 | 59.7778 | 60.648 | 59.9827 | 60.0549 | -0.08902 | -0.089 | 0.7812 | -9.8841 | -9.8119 |  | -3.2947e-09 | 2.0405e-05 | 2.0402e-05 | -2.0402e-05
13.3508 | 60.1299 | 60.0658 | 60.9189 | 60.2415 | 60.3237 | -0.064122 | -0.0641 | 0.789 | -9.8884 | -9.8062 |  | -3.4918e-09 | 2.1619e-05 | 2.1616e-05 | -2.1616e-05
13.7423 | 60.4016 | 60.3664 | 61.1922 | 60.5089 | 60.601 | -0.035223 | -0.0352 | 0.7906 | -9.8927 | -9.8006 |  | -3.7007e-09 | 2.2906e-05 | 2.2902e-05 | -2.2902e-05
14.1452 | 60.682 | 60.6795 | 61.4678 | 60.7849 | 60.8872 | -0.0025243 | -0.0025 | 0.7858 | -9.8971 | -9.7948 |  | -3.9223e-09 | 2.4269e-05 | 2.4265e-05 | -2.4265e-05
14.5599 | 60.9712 | 61.005 | 61.7455 | 61.0698 | 61.1821 | 0.033774 | 0.0338 | 0.7743 | -9.9014 | -9.7891 |  | -4.1571e-09 | 2.5713e-05 | 2.5708e-05 | -2.5708e-05
14.9868 | 61.2692 | 61.3427 | 62.0252 | 61.3635 | 61.4858 | 0.073473 | 0.0735 | 0.756 | -9.9057 | -9.7834 |  | -4.4062e-09 | 2.7243e-05 | 2.7238e-05 | -2.7238e-05
15.4263 | 61.576 | 61.6924 | 62.3067 | 61.666 | 61.7984 | 0.11637 | 0.1164 | 0.7307 | -9.91 | -9.7776 |  | -4.6703e-09 | 2.8864e-05 | 2.8859e-05 | -2.8859e-05
15.8786 | 61.8917 | 62.0537 | 62.59 | 61.9775 | 62.1197 | 0.16197 | 0.162 | 0.6983 | -9.9142 | -9.772 |  | -4.9503e-09 | 3.0581e-05 | 3.0576e-05 | -3.0576e-05
16.3442 | 62.2167 | 62.4273 | 62.875 | 62.2985 | 62.4506 | 0.21057 | 0.2106 | 0.6583 | -9.9182 | -9.7661 |  | -5.2472e-09 | 3.2401e-05 | 3.2396e-05 | -3.2396e-05
16.8234 | 62.5505 | 62.8117 | 63.1617 | 62.6285 | 62.7903 | 0.26117 | 0.2612 | 0.6112 | -9.922 | -9.7602 |  | -5.5621e-09 | 3.4329e-05 | 3.4324e-05 | -3.4324e-05
17.3167 | 62.8936 | 63.2073 | 63.4502 | 62.9679 | 63.1393 | 0.31366 | 0.3137 | 0.5566 | -9.9257 | -9.7543 |  | -5.896e-09 | 3.6372e-05 | 3.6366e-05 | -3.6366e-05
17.8244 | 63.2455 | 63.6132 | 63.7405 | 63.3165 | 63.4973 | 0.36766 | 0.3677 | 0.495 | -9.929 | -9.7482 |  | -6.2502e-09 | 3.8536e-05 | 3.853e-05 | -3.853e-05
18.347 | 63.6069 | 64.0296 | 64.033 | 63.6747 | 63.8648 | 0.42266 | 0.4227 | 0.4261 | -9.9322 | -9.7421 |  | -6.6258e-09 | 4.0829e-05 | 4.0823e-05 | -4.0823e-05
18.885 | 63.9772 | 64.4553 | 64.3279 | 64.0422 | 64.2413 | 0.47806 | 0.4781 | 0.3507 | -9.935 | -9.7359 |  | -7.0243e-09 | 4.3259e-05 | 4.3252e-05 | -4.3252e-05
19.4387 | 64.3565 | 64.8902 | 64.6256 | 64.419 | 64.627 | 0.53365 | 0.5337 | 0.2691 | -9.9375 | -9.7295 |  | -7.447e-09 | 4.5833e-05 | 4.5826e-05 | -4.5826e-05
20.0086 | 64.745 | 65.3337 | 64.9268 | 64.8054 | 65.0219 | 0.58865 | 0.5887 | 0.1818 | -9.9396 | -9.7231 |  | -7.8955e-09 | 4.856e-05 | 4.8553e-05 | -4.8553e-05
20.5953 | 65.1427 | 65.7854 | 65.232 | 65.2012 | 65.4261 | 0.64265 | 0.6427 | 0.0893 | -9.9415 | -9.7167 |  | -8.3713e-09 | 5.145e-05 | 5.1442e-05 | -5.1442e-05
21.1992 | 65.5494 | 66.2448 | 65.5424 | 65.6065 | 65.8394 | 0.69535 | 0.6954 | -0.007 | -9.9429 | -9.7101 |  | -8.8761e-09 | 5.4512e-05 | 5.4503e-05 | -5.4503e-05
21.8208 | 65.9652 | 66.7112 | 65.8587 | 66.0211 | 66.2619 | 0.74594 | 0.746 | -0.1065 | -9.9441 | -9.7034 |  | -9.4117e-09 | 5.7755e-05 | 5.7746e-05 | -5.7746e-05
22.4606 | 66.3898 | 67.184 | 66.1824 | 66.445 | 66.6932 | 0.79414 | 0.7942 | -0.2074 | -9.9448 | -9.6967 |  | -9.9802e-09 | 6.1192e-05 | 6.1182e-05 | -6.1182e-05
23.1191 | 66.8231 | 67.6625 | 66.5146 | 66.8779 | 67.1333 | 0.83934 | 0.8394 | -0.3085 | -9.9452 | -9.6899 |  | -1.0583e-08 | 6.4834e-05 | 6.4823e-05 | -6.4823e-05
23.797 | 67.2653 | 68.1466 | 66.8573 | 67.32 | 67.5823 | 0.88123 | 0.8813 | -0.408 | -9.9453 | -9.6831 |  | -1.1224e-08 | 6.8692e-05 | 6.8681e-05 | -6.8681e-05
24.4947 | 67.7159 | 68.6353 | 67.2118 | 67.7707 | 68.0395 | 0.91933 | 0.9194 | -0.5041 | -9.9452 | -9.6765 |  | -1.1904e-08 | 7.2779e-05 | 7.2767e-05 | -7.2767e-05
25.2129 | 68.1751 | 69.1286 | 67.5805 | 68.2304 | 68.5054 | 0.95342 | 0.9535 | -0.5946 | -9.9447 | -9.6698 |  | -1.2625e-08 | 7.711e-05 | 7.7098e-05 | -7.7098e-05
25.9522 | 68.6424 | 69.6255 | 67.9647 | 68.6983 | 68.9792 | 0.98302 | 0.9831 | -0.6777 | -9.9441 | -9.6633 |  | -1.3392e-08 | 8.1698e-05 | 8.1685e-05 | -8.1685e-05
26.7131 | 69.1178 | 70.1257 | 68.3665 | 69.1744 | 69.4607 | 1.0078 | 1.0079 | -0.7513 | -9.9434 | -9.6572 |  | -1.4205e-08 | 8.656e-05 | 8.6546e-05 | -8.6546e-05
27.4964 | 69.6009 | 70.6287 | 68.7877 | 69.6584 | 69.9498 | 1.0277 | 1.0278 | -0.8132 | -9.9425 | -9.6512 |  | -1.507e-08 | 9.1711e-05 | 9.1696e-05 | -9.1696e-05
28.3026 | 70.0915 | 71.1341 | 69.2298 | 70.15 | 70.446 | 1.0425 | 1.0426 | -0.8617 | -9.9415 | -9.6456 |  | -1.5987e-08 | 9.7169e-05 | 9.7153e-05 | -9.7153e-05
29.1325 | 70.5894 | 71.6414 | 69.6939 | 70.6487 | 70.9488 | 1.0519 | 1.052 | -0.8955 | -9.9407 | -9.6407 |  | -1.6963e-08 | 0.00010295 | 0.00010293 | -0.00010293
29.9866 | 71.0939 | 72.1501 | 70.1805 | 71.1539 | 71.4578 | 1.0561 | 1.0562 | -0.9134 | -9.94 | -9.6362 |  | -1.7999e-08 | 0.00010908 | 0.00010906 | -0.00010906
30.8659 | 71.605 | 72.6601 | 70.6905 | 71.6655 | 71.9726 | 1.055 | 1.0551 | -0.9145 | -9.9395 | -9.6325 |  | -1.91e-08 | 0.00011557 | 0.00011555 | -0.00011555
31.7709 | 72.122 | 73.1708 | 71.2225 | 72.1827 | 72.4925 | 1.0487 | 1.0488 | -0.8995 | -9.9393 | -9.6296 |  | -2.027e-08 | 0.00012244 | 0.00012242 | -0.00012242
32.7024 | 72.6446 | 73.6821 | 71.7763 | 72.7051 | 73.017 | 1.0374 | 1.0375 | -0.8683 | -9.9395 | -9.6277 |  | -2.1514e-08 | 0.00012973 | 0.00012971 | -0.00012971
33.6613 | 73.1721 | 74.1934 | 72.3499 | 73.2318 | 73.5454 | 1.0212 | 1.0213 | -0.8222 | -9.9403 | -9.6268 |  | -2.2837e-08 | 0.00013745 | 0.00013743 | -0.00013743
34.6483 | 73.704 | 74.7047 | 72.9409 | 73.7624 | 74.0769 | 1.0006 | 1.0007 | -0.7631 | -9.9416 | -9.6272 |  | -2.4243e-08 | 0.00014563 | 0.00014561 | -0.00014561
35.6642 | 74.2396 | 75.2155 | 73.5468 | 74.2959 | 74.6107 | 0.97575 | 0.9759 | -0.6928 | -9.9437 | -9.6291 |  | -2.5739e-08 | 0.0001543 | 0.00015427 | -0.00015427
36.7099 | 74.7783 | 75.7258 | 74.1645 | 74.8317 | 75.1462 | 0.94734 | 0.9475 | -0.6138 | -9.9466 | -9.6323 |  | -2.7331e-08 | 0.00016348 | 0.00016345 | -0.00016345
37.7862 | 75.3194 | 76.2351 | 74.7906 | 75.3688 | 75.6823 | 0.91553 | 0.9157 | -0.5288 | -9.9506 | -9.6373 |  | -2.9025e-08 | 0.00017321 | 0.00017318 | -0.00017318
38.8942 | 75.8621 | 76.7433 | 75.4216 | 75.9063 | 76.2182 | 0.88102 | 0.8812 | -0.4405 | -9.9558 | -9.6441 |  | -3.0828e-08 | 0.00018351 | 0.00018348 | -0.00018348
40.0346 | 76.4056 | 77.2501 | 76.0538 | 76.4432 | 76.7527 | 0.84431 | 0.8445 | -0.3518 | -9.9624 | -9.6531 |  | -3.2747e-08 | 0.00019443 | 0.0001944 | -0.0001944
41.2084 | 76.9489 | 77.7553 | 76.6837 | 76.9786 | 77.285 | 0.80619 | 0.8064 | -0.2652 | -9.9703 | -9.6641 |  | -3.4791e-08 | 0.000206 | 0.00020597 | -0.00020597
42.4167 | 77.4913 | 78.2584 | 77.308 | 77.5112 | 77.8139 | 0.76688 | 0.7671 | -0.1833 | -9.9801 | -9.6776 |  | -3.6968e-08 | 0.00021826 | 0.00021823 | -0.00021823
43.6604 | 78.0316 | 78.7592 | 77.9233 | 78.0399 | 78.3382 | 0.72737 | 0.7276 | -0.1083 | -9.9917 | -9.6936 |  | -3.9288e-08 | 0.00023125 | 0.00023121 | -0.00023121
44.9405 | 78.569 | 79.2572 | 78.5265 | 78.5634 | 78.8568 | 0.68796 | 0.6882 | -0.0425 | -10.0056 | -9.7124 |  | -4.1761e-08 | 0.00024501 | 0.00024497 | -0.00024497
46.2582 | 79.1023 | 79.752 | 79.1149 | 79.0804 | 79.3682 | 0.64944 | 0.6497 | 0.0126 | -10.0219 | -9.7344 |  | -4.4397e-08 | 0.00025959 | 0.00025955 | -0.00025955
47.6145 | 79.6305 | 80.2428 | 79.6859 | 79.5896 | 79.8714 | 0.61203 | 0.6123 | 0.0554 | -10.0409 | -9.7594 |  | -4.7209e-08 | 0.00027504 | 0.000275 | -0.000275
49.0106 | 80.1525 | 80.7291 | 80.2373 | 80.0893 | 80.3647 | 0.57631 | 0.5766 | 0.0848 | -10.0632 | -9.7881 |  | -5.0209e-08 | 0.00029141 | 0.00029136 | -0.00029136
50.4477 | 80.6671 | 81.21 | 80.7674 | 80.5781 | 80.8468 | 0.54259 | 0.5429 | 0.1003 | -10.089 | -9.8206 |  | -5.3411e-08 | 0.00030875 | 0.0003087 | -0.0003087
51.9268 | 81.1731 | 81.6845 | 81.2746 | 81.0543 | 81.3162 | 0.51107 | 0.5114 | 0.1015 | -10.1188 | -9.8572 |  | -5.683e-08 | 0.00032713 | 0.00032707 | -0.00032707
53.4494 | 81.6694 | 82.1518 | 81.7581 | 81.5163 | 81.7714 | 0.48205 | 0.4824 | 0.0887 | -10.1531 | -9.8983 |  | -6.0481e-08 | 0.00034659 | 0.00034653 | -0.00034653
55.0165 | 82.1547 | 82.6106 | 82.2174 | 81.9621 | 82.2107 | 0.45553 | 0.4559 | 0.0627 | -10.1926 | -9.9444 |  | -6.4384e-08 | 0.00036722 | 0.00036716 | -0.00036716
56.6297 | 82.628 | 83.0597 | 82.6528 | 82.3901 | 82.6325 | 0.43131 | 0.4317 | 0.0248 | -10.2379 | -9.9959 |  | -6.8555e-08 | 0.00038907 | 0.00038901 | -0.00038901
58.2901 | 83.0881 | 83.4981 | 83.0652 | 82.7983 | 83.0351 | 0.40959 | 0.41 | -0.0229 | -10.2898 | -10.0534 |  | -7.3016e-08 | 0.00041223 | 0.00041216 | -0.00041216
59.9992 | 83.5339 | 83.9244 | 83.4562 | 83.1848 | 83.417 | 0.39006 | 0.3905 | -0.0777 | -10.3491 | -10.1173 |  | -7.779e-08 | 0.00043676 | 0.00043669 | -0.00043669
61.7584 | 83.9646 | 84.3375 | 83.8282 | 83.5478 | 83.7766 | 0.37244 | 0.3729 | -0.1364 | -10.4168 | -10.1885 |  | -8.29e-08 | 0.00046276 | 0.00046267 | -0.00046267
63.5692 | 84.3795 | 84.7363 | 84.1846 | 83.8855 | 84.1123 | 0.35631 | 0.3568 | -0.1949 | -10.494 | -10.2677 |  | -8.8373e-08 | 0.0004903 | 0.00049021 | -0.00049021
65.4331 | 84.7779 | 85.1203 | 84.5291 | 84.1964 | 84.423 | 0.34188 | 0.3424 | -0.2488 | -10.5815 | -10.3554 |  | -9.4238e-08 | 0.00051948 | 0.00051938 | -0.00051938
67.3517 | 85.1599 | 85.4888 | 84.8662 | 84.4791 | 84.7076 | 0.32835 | 0.3289 | -0.2937 | -10.6808 | -10.4529 |  | -1.0053e-07 | 0.00055039 | 0.00055029 | -0.00055029
69.3265 | 85.5256 | 85.8421 | 85.2002 | 84.733 | 84.9658 | 0.31592 | 0.3165 | -0.3254 | -10.7926 | -10.5604 |  | -1.0727e-07 | 0.00058315 | 0.00058304 | -0.00058304
71.3592 | 85.8759 | 86.1807 | 85.5353 | 84.9578 | 85.1975 | 0.30418 | 0.3048 | -0.3406 | -10.9181 | -10.679 |  | -1.1451e-07 | 0.00061786 | 0.00061774 | -0.00061774
73.4515 | 86.2122 | 86.5061 | 85.8748 | 85.1546 | 85.4041 | 0.29325 | 0.2939 | -0.3374 | -11.0576 | -10.8088 |  | -1.2229e-07 | 0.00065463 | 0.00065451 | -0.00065451
75.6051 | 86.5368 | 86.8205 | 86.2211 | 85.3258 | 85.5877 | 0.28301 | 0.2837 | -0.3157 | -11.211 | -10.9498 |  | -1.3065e-07 | 0.00069359 | 0.00069346 | -0.00069346
77.8219 | 86.8526 | 87.127 | 86.5748 | 85.4754 | 85.7524 | 0.27367 | 0.2744 | -0.2778 | -11.3772 | -11.1009 |  | -1.3964e-07 | 0.00073487 | 0.00073473 | -0.00073473
80.1037 | 87.1635 | 87.4298 | 86.9353 | 85.61 | 85.904 | 0.26552 | 0.2663 | -0.2282 | -11.5535 | -11.2603 |  | -1.4931e-07 | 0.00077861 | 0.00077846 | -0.00077846
82.4524 | 87.4741 | 87.7335 | 87.3012 | 85.7387 | 86.0506 | 0.25858 | 0.2594 | -0.1729 | -11.7354 | -11.4243 |  | -1.5972e-07 | 0.00082496 | 0.0008248 | -0.0008248
84.87 | 87.7895 | 88.0436 | 87.6705 | 85.8731 | 86.2024 | 0.25323 | 0.2541 | -0.119 | -11.9164 | -11.588 |  | -1.7094e-07 | 0.00087406 | 0.00087389 | -0.00087389
87.3584 | 88.1148 | 88.3652 | 88.0419 | 86.0273 | 86.3711 | 0.24947 | 0.2504 | -0.0729 | -12.0875 | -11.7446 |  | -1.8304e-07 | 0.00092608 | 0.0009259 | -0.0009259
89.9199 | 88.4545 | 88.7029 | 88.4151 | 86.2159 | 86.5691 | 0.24742 | 0.2484 | -0.0394 | -12.2386 | -11.8864 |  | -1.961e-07 | 0.00098121 | 0.00098101 | -0.00098101
92.5564 | 88.812 | 89.0597 | 88.7918 | 86.4522 | 86.8078 | 0.24666 | 0.2477 | -0.0202 | -12.3598 | -12.0052 |  | -2.1019e-07 | 0.0010396 | 0.0010394 | -0.0010394
95.2702 | 89.1891 | 89.4369 | 89.1754 | 86.7456 | 87.0951 | 0.2467 | 0.2478 | -0.0137 | -12.4435 | -12.0951 |  | -2.2543e-07 | 0.0011015 | 0.0011013 | -0.0011013
98.0636 | 89.5849 | 89.8333 | 89.5705 | 87.0987 | 87.4341 | 0.24723 | 0.2484 | -0.0144 | -12.4862 | -12.152 |  | -2.419e-07 | 0.0011671 | 0.0011668 | -0.0011668
100.9389 | 89.9966 | 90.2451 | 89.9819 | 87.5069 | 87.8213 | 0.24726 | 0.2485 | -0.0147 | -12.4897 | -12.1765 |  | -2.5973e-07 | 0.0012365 | 0.0012363 | -0.0012363
103.8985 | 90.4191 | 90.6668 | 90.4129 | 87.9586 | 88.2481 | 0.24639 | 0.2477 | -0.0062 | -12.4605 | -12.1723 |  | -2.7904e-07 | 0.0013101 | 0.0013099 | -0.0013099
106.9449 | 90.846 | 91.0914 | 90.8636 | 88.4378 | 88.7015 | 0.24401 | 0.2454 | 0.0176 | -12.4082 | -12.1459 |  | -2.9997e-07 | 0.0013881 | 0.0013878 | -0.0013878
110.0806 | 91.2702 | 91.5116 | 91.3301 | 88.9274 | 89.1669 | 0.23993 | 0.2414 | 0.0599 | -12.3428 | -12.1048 |  | -3.2268e-07 | 0.0014708 | 0.0014705 | -0.0014705
113.3082 | 91.6851 | 91.9207 | 91.8046 | 89.4114 | 89.6305 | 0.23404 | 0.2356 | 0.1195 | -12.2737 | -12.0562 |  | -3.4734e-07 | 0.0015583 | 0.001558 | -0.001558
116.6305 | 92.0851 | 92.3134 | 92.2764 | 89.8777 | 90.081 | 0.22665 | 0.2283 | 0.1913 | -12.2074 | -12.0058 |  | -3.7413e-07 | 0.0016511 | 0.0016507 | -0.0016507
120.0502 | 92.4668 | 92.6863 | 92.7335 | 90.3184 | 90.5107 | 0.21775 | 0.2195 | 0.2667 | -12.1484 | -11.9578 |  | -4.0326e-07 | 0.0017494 | 0.001749 | -0.001749
123.5702 | 92.8288 | 93.0383 | 93.1654 | 90.7309 | 90.9165 | 0.20765 | 0.2095 | 0.3366 | -12.0979 | -11.9142 |  | -4.3498e-07 | 0.0018536 | 0.0018531 | -0.0018531
127.1933 | 93.1724 | 93.3712 | 93.5641 | 91.1172 | 91.299 | 0.19684 | 0.1988 | 0.3917 | -12.0552 | -11.8754 |  | -4.6952e-07 | 0.0019639 | 0.0019635 | -0.0019635
130.9228 | 93.5009 | 93.6884 | 93.9258 | 91.4831 | 91.6628 | 0.18542 | 0.1875 | 0.4249 | -12.0178 | -11.8402 |  | -5.0718e-07 | 0.0020809 | 0.0020804 | -0.0020804
134.7615 | 93.8189 | 93.9952 | 94.2512 | 91.837 | 92.0145 | 0.1741 | 0.1763 | 0.4323 | -11.9819 | -11.8066 |  | -5.4828e-07 | 0.0022048 | 0.0022042 | -0.0022042
138.7128 | 94.1317 | 94.2968 | 94.5456 | 92.1877 | 92.3618 | 0.16276 | 0.1651 | 0.4139 | -11.944 | -11.7722 |  | -5.9316e-07 | 0.0023361 | 0.0023355 | -0.0023355
142.78 | 94.4436 | 94.5977 | 94.8178 | 92.542 | 92.711 | 0.15163 | 0.1541 | 0.3742 | -11.9016 | -11.7351 |  | -6.4221e-07 | 0.0024752 | 0.0024746 | -0.0024746
146.9664 | 94.7569 | 94.9004 | 95.0784 | 92.9037 | 93.0656 | 0.14088 | 0.1435 | 0.3215 | -11.8532 | -11.6939 |  | -6.9586e-07 | 0.0026226 | 0.0026219 | -0.0026219
151.2756 | 95.0712 | 95.2045 | 95.3379 | 93.2717 | 93.425 | 0.13052 | 0.1333 | 0.2667 | -11.7995 | -11.649 |  | -7.546e-07 | 0.0027788 | 0.002778 | -0.002778
155.7111 | 95.3833 | 95.5065 | 95.6036 | 93.6415 | 93.785 | 0.12026 | 0.1232 | 0.2203 | -11.7418 | -11.6012 |  | -8.1896e-07 | 0.0029443 | 0.0029435 | -0.0029435
160.2766 | 95.6879 | 95.8014 | 95.878 | 94.0057 | 94.1389 | 0.11038 | 0.1135 | 0.1901 | -11.6822 | -11.5521 |  | -8.8954e-07 | 0.0031197 | 0.0031188 | -0.0031188
164.9761 | 95.9788 | 96.0825 | 96.1575 | 94.3564 | 94.4788 | 0.1004 | 0.1037 | 0.1787 | -11.6224 | -11.5033 |  | -9.6699e-07 | 0.0033055 | 0.0033045 | -0.0033045
169.8133 | 96.2496 | 96.3434 | 96.4331 | 94.6865 | 94.7974 | 0.090299 | 0.0938 | 0.1835 | -11.5631 | -11.4557 |  | -1.0521e-06 | 0.0035024 | 0.0035013 | -0.0035013
174.7924 | 96.4941 | 96.5779 | 96.6919 | 94.9899 | 95.0887 | 0.08009 | 0.0838 | 0.1978 | -11.5042 | -11.4091 |  | -1.1456e-06 | 0.0037111 | 0.0037099 | -0.0037099
179.9174 | 96.7068 | 96.7803 | 96.9198 | 95.2619 | 95.3478 | 0.069569 | 0.0735 | 0.213 | -11.4449 | -11.3629 |  | -1.2484e-06 | 0.0039321 | 0.0039309 | -0.0039309
185.1927 | 96.8816 | 96.9446 | 97.1034 | 95.4975 | 95.5703 | 0.058835 | 0.063 | 0.2218 | -11.3841 | -11.3155 |  | -1.3616e-06 | 0.0041664 | 0.0041651 | -0.0041651
190.6227 | 97.0119 | 97.064 | 97.232 | 95.6914 | 95.7511 | 0.047687 | 0.0521 | 0.2201 | -11.3205 | -11.2652 |  | -1.4864e-06 | 0.0044147 | 0.0044132 | -0.0044132
196.2119 | 97.0901 | 97.1309 | 97.2976 | 95.8365 | 95.8837 | 0.036124 | 0.0408 | 0.2075 | -11.2536 | -11.2111 |  | -1.6239e-06 | 0.0046778 | 0.0046762 | -0.0046762
201.965 | 97.1087 | 97.1379 | 97.2953 | 95.9248 | 95.9604 | 0.024245 | 0.0292 | 0.1866 | -11.1839 | -11.1533 |  | -1.7756e-06 | 0.0049566 | 0.0049548 | -0.0049548
207.8868 | 97.0617 | 97.0782 | 97.2223 | 95.9477 | 95.9731 | 0.01125 | 0.0165 | 0.1606 | -11.114 | -11.0939 |  | -1.9431e-06 | 0.005252 | 0.00525 | -0.00525
213.9822 | 96.947 | 96.9487 | 97.0786 | 95.8982 | 95.9154 | -0.0038629 | 0.0017 | 0.1316 | -11.0488 | -11.0372 |  | -2.1281e-06 | 0.005565 | 0.0055629 | -0.0055629
220.2563 | 96.7683 | 96.7517 | 96.8682 | 95.7723 | 95.7848 | -0.022494 | -0.0166 | 0.0999 | -10.996 | -10.9894 |  | -2.3327e-06 | 0.0058968 | 0.0058945 | -0.0058945
226.7144 | 96.5363 | 96.497 | 96.6013 | 95.5733 | 95.5852 | -0.045546 | -0.0393 | 0.065 | -10.963 | -10.9573 |  | -2.559e-06 | 0.0062484 | 0.0062458 | -0.0062458
233.3618 | 96.2681 | 96.2035 | 96.2964 | 95.3135 | 95.3288 | -0.071218 | -0.0646 | 0.0283 | -10.9546 | -10.9459 |  | -2.8095e-06 | 0.0066209 | 0.0066181 | -0.0066181
240.2041 | 95.9854 | 95.8987 | 95.9804 | 95.0167 | 95.0376 | -0.093713 | -0.0867 | -0.005 | -10.9687 | -10.9548 |  | -3.0869e-06 | 0.0070158 | 0.0070127 | -0.0070127
247.2471 | 95.7109 | 95.6155 | 95.6851 | 94.7158 | 94.7414 | -0.10283 | -0.0954 | -0.0258 | -10.9951 | -10.9769 |  | -3.3943e-06 | 0.0074342 | 0.0074308 | -0.0074308
254.4966 | 95.4679 | 95.386 | 95.4423 | 94.4468 | 94.4732 | -0.089774 | -0.0819 | -0.0256 | -11.0211 | -11.0026 |  | -3.7352e-06 | 0.0078777 | 0.0078739 | -0.0078739
261.9586 | 95.2812 | 95.2348 | 95.2801 | 94.2395 | 94.2628 | -0.054744 | -0.0464 | -0.0011 | -11.0417 | -11.0267 |  | -4.1135e-06 | 0.0083477 | 0.0083435 | -0.0083435
269.6394 | 95.1757 | 95.1722 | 95.2173 | 94.1096 | 94.1277 | -0.012341 | -0.0035 | 0.0416 | -11.0661 | -11.0568 |  | -4.5333e-06 | 0.0088457 | 0.0088412 | -0.0088412
277.5455 | 95.1638 | 95.1883 | 95.2518 | 94.0531 | 94.0655 | 0.015131 | 0.0245 | 0.088 | -11.1107 | -11.1077 |  | -4.9997e-06 | 0.0093736 | 0.0093686 | -0.0093686
285.6833 | 95.2268 | 95.2511 | 95.3497 | 94.0437 | 94.0498 | 0.014372 | 0.0243 | 0.1229 | -11.1831 | -11.1869 |  | -5.5178e-06 | 0.0099331 | 0.0099276 | -0.0099276
294.0598 | 95.3116 | 95.3117 | 95.4506 | 94.0378 | 94.0358 | -0.01042 | 0.0001 | 0.139 | -11.2738 | -11.2863 |  | -6.0939e-06 | 0.010526 | 0.01052 | -0.01052
302.6818 | 95.3588 | 95.3248 | 95.4968 | 93.9913 | 93.9804 | -0.045148 | -0.034 | 0.138 | -11.3675 | -11.3895 |  | -6.7346e-06 | 0.011155 | 0.011148 | -0.011148
311.5567 | 95.3437 | 95.2763 | 95.4689 | 93.8864 | 93.8686 | -0.079213 | -0.0674 | 0.1252 | -11.4573 | -11.4869 |  | -7.4476e-06 | 0.011821 | 0.011813 | -0.011813
320.6918 | 95.295 | 95.1975 | 95.402 | 93.7487 | 93.7283 | -0.11002 | -0.0975 | 0.107 | -11.5463 | -11.5792 |  | -8.2412e-06 | 0.012527 | 0.012519 | -0.012519
330.0947 | 95.2672 | 95.1417 | 95.3575 | 93.6341 | 93.6122 | -0.13877 | -0.1255 | 0.0903 | -11.6331 | -11.6683 |  | -9.1251e-06 | 0.013276 | 0.013266 | -0.013266
339.7733 | 95.2795 | 95.1296 | 95.3599 | 93.5812 | 93.5514 | -0.16396 | -0.1499 | 0.0804 | -11.6983 | -11.7422 |  | -1.011e-05 | 0.014069 | 0.014059 | -0.014059
349.7357 | 95.2908 | 95.123 | 95.3637 | 93.5751 | 93.5278 | -0.1827 | -0.1678 | 0.0729 | -11.7157 | -11.7779 |  | -1.1207e-05 | 0.01491 | 0.014899 | -0.014899
359.9902 | 95.2478 | 95.0708 | 95.3035 | 93.57 | 93.503 | -0.19279 | -0.177 | 0.0557 | -11.6778 | -11.7606 |  | -1.2431e-05 | 0.015802 | 0.015789 | -0.015789
370.5454 | 95.1497 | 94.976 | 95.1737 | 93.5412 | 93.4651 | -0.19043 | -0.1737 | 0.024 | -11.6085 | -11.7013 |  | -1.3796e-05 | 0.016747 | 0.016733 | -0.016733
381.4101 | 95.0607 | 94.9067 | 95.0462 | 93.5049 | 93.4391 | -0.17173 | -0.154 | -0.0145 | -11.5558 | -11.6393 |  | -1.5318e-05 | 0.017749 | 0.017733 | -0.017733
392.5933 | 95.0633 | 94.9395 | 95.0163 | 93.5073 | 93.4665 | -0.14259 | -0.1238 | -0.047 | -11.556 | -11.6156 |  | -1.7018e-05 | 0.018811 | 0.018794 | -0.018794
404.1045 | 95.1997 | 95.0991 | 95.1339 | 93.6075 | 93.5887 | -0.12052 | -0.1006 | -0.0658 | -11.5922 | -11.6309 |  | -1.8915e-05 | 0.019937 | 0.019918 | -0.019918
415.9531 | 95.4474 | 95.3493 | 95.3715 | 93.8354 | 93.8224 | -0.11921 | -0.0981 | -0.0759 | -11.612 | -11.6461 |  | -2.1033e-05 | 0.021131 | 0.02111 | -0.02111
428.1492 | 95.728 | 95.6161 | 95.6373 | 94.1454 | 94.1259 | -0.13427 | -0.1119 | -0.0907 | -11.5826 | -11.6245 |  | -2.34e-05 | 0.022398 | 0.022374 | -0.022374
440.7028 | 95.9419 | 95.8144 | 95.8372 | 94.4571 | 94.4238 | -0.15121 | -0.1275 | -0.1047 | -11.4848 | -11.5418 |  | -2.6043e-05 | 0.02374 | 0.023714 | -0.023714
453.6246 | 96.0245 | 95.887 | 95.9318 | 94.7191 | 94.6655 | -0.16263 | -0.1375 | -0.0927 | -11.3054 | -11.3841 |  | -2.8998e-05 | 0.025164 | 0.025135 | -0.025135
466.9252 | 95.9793 | 95.8373 | 95.9272 | 94.8926 | 94.8231 | -0.16864 | -0.142 | -0.0521 | -11.0867 | -11.1828 |  | -3.2302e-05 | 0.026674 | 0.026641 | -0.026641
480.6158 | 95.8616 | 95.7183 | 95.841 | 94.9421 | 94.8693 | -0.17154 | -0.1433 | -0.0206 | -10.9195 | -11.0205 |  | -3.5995e-05 | 0.028275 | 0.028239 | -0.028239
494.7078 | 95.7465 | 95.5926 | 95.7055 | 94.8992 | 94.815 | -0.18383 | -0.1539 | -0.041 | -10.8473 | -10.9614 |  | -4.0125e-05 | 0.029973 | 0.029933 | -0.029933
509.213 | 95.6772 | 95.4961 | 95.5676 | 94.8461 | 94.7173 | -0.21283 | -0.1811 | -0.1096 | -10.8311 | -10.9916 |  | -4.4746e-05 | 0.031774 | 0.031729 | -0.031729
524.1435 | 95.602 | 95.4015 | 95.4314 | 94.7716 | 94.5963 | -0.23413 | -0.2005 | -0.1706 | -10.8304 | -11.0393 |  | -4.9916e-05 | 0.033684 | 0.033634 | -0.033634
539.5118 | 95.447 | 95.2524 | 95.2621 | 94.5812 | 94.4255 | -0.23025 | -0.1946 | -0.1849 | -10.8658 | -11.0572 |  | -5.5701e-05 | 0.035711 | 0.035655 | -0.035655
555.3307 | 95.2481 | 95.0627 | 95.0867 | 94.3144 | 94.2247 | -0.2232 | -0.1854 | -0.1614 | -10.9337 | -11.0612 |  | -6.2176e-05 | 0.03786 | 0.037798 | -0.037798
571.6134 | 95.1059 | 94.9191 | 94.983 | 94.1324 | 94.0593 | -0.22687 | -0.1868 | -0.1229 | -10.9735 | -11.0867 |  | -6.9424e-05 | 0.040141 | 0.040072 | -0.040072
588.3736 | 95.0357 | 94.8386 | 94.9389 | 94.0397 | 93.9408 | -0.23958 | -0.1971 | -0.0968 | -10.996 | -11.1374 |  | -7.7539e-05 | 0.042561 | 0.042484 | -0.042484
605.6251 | 94.9258 | 94.7218 | 94.8444 | 93.8931 | 93.779 | -0.24904 | -0.204 | -0.0814 | -11.0327 | -11.1918 |  | -8.6628e-05 | 0.045129 | 0.045042 | -0.045042
623.3825 | 94.7097 | 94.512 | 94.6794 | 93.616 | 93.492 | -0.24546 | -0.1977 | -0.0303 | -11.0937 | -11.2655 |  | -9.6806e-05 | 0.047854 | 0.047757 | -0.047757
641.6606 | 94.5514 | 94.3787 | 94.6058 | 93.3687 | 93.2323 | -0.22334 | -0.1727 | 0.0544 | -11.1827 | -11.3697 |  | -0.00010821 | 0.050746 | 0.050637 | -0.050637
660.4746 | 94.6375 | 94.4637 | 94.7707 | 93.3738 | 93.2128 | -0.22749 | -0.1738 | 0.1332 | -11.2637 | -11.4784 |  | -0.00012098 | 0.053815 | 0.053694 | -0.053694
679.8402 | 94.7074 | 94.5656 | 94.8378 | 93.3581 | 93.176 | -0.19874 | -0.1418 | 0.1304 | -11.3493 | -11.5883 |  | -0.00013529 | 0.057073 | 0.056938 | -0.056938
699.7736 | 94.4522 | 94.2774 | 94.6306 | 93.0375 | 92.8452 | -0.23518 | -0.1748 | 0.1784 | -11.4147 | -11.6674 |  | -0.00015133 | 0.060532 | 0.060381 | -0.060381
720.2915 | 94.0579 | 93.4256 | 93.8812 | 92.639 | 92.4113 | -0.69634 | -0.6323 | -0.1767 | -11.4189 | -11.7106 |  | -0.00016931 | 0.064205 | 0.064036 | -0.064036
741.411 | 93.838 | 93.1583 | 93.3594 | 92.3852 | 92.1585 | -0.74762 | -0.6797 | -0.4786 | -11.4528 | -11.7474 |  | -0.00018945 | 0.068105 | 0.067916 | -0.067916
763.1497 | 93.7053 | 93.3337 | 93.4378 | 92.2039 | 91.9906 | -0.44363 | -0.3716 | -0.2675 | -11.5014 | -11.7867 |  | -0.00021204 | 0.072247 | 0.072035 | -0.072035
785.5259 | 93.5046 | 93.3628 | 93.3852 | 91.842 | 91.6725 | -0.21821 | -0.1418 | -0.1194 | -11.6626 | -11.9085 |  | -0.00023736 | 0.076646 | 0.076409 | -0.076409
808.5581 | 93.897 | 93.7006 | 93.7845 | 92.0642 | 91.9641 | -0.27745 | -0.1964 | -0.1125 | -11.8328 | -12.014 |  | -0.00026575 | 0.08132 | 0.081054 | -0.081054
832.2656 | 94.8877 | 94.6219 | 94.7108 | 93.1485 | 93.0643 | -0.35179 | -0.2658 | -0.1769 | -11.7392 | -11.9094 |  | -0.00029759 | 0.086286 | 0.085988 | -0.085988
856.6683 | 95.5085 | 95.2391 | 95.3344 | 94.0305 | 93.9375 | -0.36063 | -0.2694 | -0.1741 | -11.478 | -11.6622 |  | -0.0003333 | 0.091562 | 0.091229 | -0.091229
881.7864 | 95.4188 | 95.1197 | 95.2319 | 94.2288 | 94.1163 | -0.3959 | -0.2991 | -0.1869 | -11.19 | -11.3993 |  | -0.00037334 | 0.097171 | 0.096798 | -0.096798
907.6411 | 94.9325 | 94.6192 | 94.7425 | 93.9432 | 93.7948 | -0.41601 | -0.3133 | -0.19 | -10.9893 | -11.2404 |  | -0.00041826 | 0.10313 | 0.10271 | -0.10271
934.2538 | 94.3735 | 94.0781 | 94.2095 | 93.4156 | 93.234 | -0.4044 | -0.2954 | -0.164 | -10.9579 | -11.2485 |  | -0.00046866 | 0.10947 | 0.109 | -0.109
961.6468 | 94.0345 | 93.7911 | 93.8825 | 92.9144 | 92.726 | -0.35909 | -0.2434 | -0.152 | -11.1201 | -11.4242 |  | -0.00052519 | 0.11621 | 0.11569 | -0.11569
989.843 | 94.042 | 93.8451 | 93.8692 | 92.6149 | 92.4599 | -0.3197 | -0.1969 | -0.1728 | -11.4271 | -11.7049 |  | -0.00058862 | 0.12338 | 0.1228 | -0.1228
1018.866 | 94.146 | 93.9673 | 93.9262 | 92.5607 | 92.4446 | -0.30905 | -0.1787 | -0.2198 | -11.5853 | -11.8318 |  | -0.0006598 | 0.13101 | 0.13035 | -0.13035
1048.7399 | 93.9976 | 93.816 | 93.7608 | 92.5351 | 92.4278 | -0.31999 | -0.1816 | -0.2368 | -11.4625 | -11.7082 |  | -0.00073967 | 0.13913 | 0.13839 | -0.13839
1079.4898 | 93.7697 | 93.5817 | 93.589 | 92.4349 | 92.3137 | -0.33495 | -0.188 | -0.1807 | -11.3348 | -11.6029 |  | -0.00082932 | 0.14778 | 0.14695 | -0.14695
1111.1412 | 93.7609 | 93.558 | 93.6843 | 92.3496 | 92.2097 | -0.35895 | -0.2029 | -0.0766 | -11.4113 | -11.7072 |  | -0.00092993 | 0.15698 | 0.15605 | -0.15605
1143.7207 | 93.9732 | 93.7757 | 93.9311 | 92.4002 | 92.2519 | -0.36324 | -0.1975 | -0.0421 | -11.573 | -11.887 |  | -0.0010429 | 0.16678 | 0.16574 | -0.16574
1177.2555 | 94.2044 | 94.0094 | 94.0739 | 92.6133 | 92.4754 | -0.37105 | -0.195 | -0.1305 | -11.5911 | -11.9051 |  | -0.0011696 | 0.17722 | 0.17605 | -0.17605
1211.7735 | 94.2255 | 94.034 | 93.9521 | 92.8305 | 92.6945 | -0.37854 | -0.1915 | -0.2734 | -11.395 | -11.718 |  | -0.001312 | 0.18835 | 0.18704 | -0.18704
1247.3036 | 94.0424 | 93.8953 | 93.6294 | 92.8443 | 92.728 | -0.34584 | -0.1471 | -0.413 | -11.1981 | -11.5131 |  | -0.0014718 | 0.20021 | 0.19874 | -0.19874
1283.8755 | 93.9111 | 93.8202 | 93.4626 | 92.741 | 92.6986 | -0.30211 | -0.0909 | -0.4485 | -11.1701 | -11.4237 |  | -0.0016512 | 0.21286 | 0.21121 | -0.21121
1321.5197 | 93.8784 | 93.8024 | 93.5241 | 92.7574 | 92.7814 | -0.30051 | -0.076 | -0.3543 | -11.121 | -11.3215 |  | -0.0018526 | 0.22636 | 0.22451 | -0.22451
1360.2677 | 93.8967 | 93.8081 | 93.6243 | 92.8222 | 92.8412 | -0.32729 | -0.0886 | -0.2724 | -11.0745 | -11.2942 |  | -0.0020789 | 0.24077 | 0.23869 | -0.23869
1400.1518 | 94.0105 | 93.9145 | 93.6904 | 92.7164 | 92.702 | -0.34981 | -0.096 | -0.3201 | -11.2941 | -11.5623 |  | -0.002333 | 0.25614 | 0.25381 | -0.25381
1441.2053 | 94.4259 | 94.2913 | 93.9401 | 92.948 | 92.9336 | -0.40455 | -0.1346 | -0.4858 | -11.4779 | -11.7623 |  | -0.0026185 | 0.27257 | 0.26995 | -0.26995
1483.4625 | 95.0732 | 94.9632 | 94.5088 | 93.8267 | 93.8335 | -0.39719 | -0.11 | -0.5644 | -11.2465 | -11.5269 |  | -0.0029391 | 0.29013 | 0.28719 | -0.28719
1526.9588 | 95.0093 | 95.0115 | 94.4847 | 93.9599 | 94.0153 | -0.30339 | 0.0022 | -0.5246 | -11.0494 | -11.2996 |  | -0.0032993 | 0.30889 | 0.30559 | -0.30559
1571.7304 | 94.9793 | 95.0957 | 94.5976 | 93.8696 | 93.9803 | -0.20886 | 0.1164 | -0.3817 | -11.1097 | -11.3243 |  | -0.0037039 | 0.32897 | 0.32526 | -0.32526
1617.8147 | 95.4644 | 95.6125 | 95.1894 | 93.9665 | 94.097 | -0.19819 | 0.1481 | -0.275 | -11.4979 | -11.7137 |  | -0.0041586 | 0.35045 | 0.34629 | -0.34629
1665.2502 | 96.0432 | 96.1124 | 95.5677 | 94.2548 | 94.4321 | -0.29958 | 0.0692 | -0.4755 | -11.7884 | -11.9799 |  | -0.0046695 | 0.37345 | 0.36878 | -0.36878
1714.0766 | 95.5948 | 95.5672 | 94.9667 | 94.3681 | 94.5644 | -0.42045 | -0.0276 | -0.6281 | -11.2267 | -11.4233 |  | -0.0052437 | 0.3981 | 0.39285 | -0.39285
1764.3347 | 95.9351 | 95.9289 | 95.3502 | 95.4539 | 95.6049 | -0.42483 | -0.0062 | -0.5849 | -10.4812 | -10.7488 |  | -0.005889 | 0.42452 | 0.41863 | -0.41863
1816.0663 | 95.4958 | 95.6439 | 95.1564 | 95.1283 | 95.2429 | -0.29814 | 0.1481 | -0.3394 | -10.3675 | -10.6991 |  | -0.0066144 | 0.45286 | 0.44624 | -0.44624
1869.3148 | 95.5687 | 95.9298 | 95.2626 | 94.7186 | 94.8991 | -0.11474 | 0.3611 | -0.3061 | -10.8501 | -11.1454 |  | -0.0074298 | 0.48327 | 0.47584 | -0.47584
1924.1245 | 95.7428 | 96.1243 | 95.3642 | 94.5653 | 94.8196 | -0.1261 | 0.3815 | -0.3786 | -11.1775 | -11.4308 |  | -0.0083465 | 0.51594 | 0.5076 | -0.5076
1980.5413 | 95.9985 | 96.3494 | 95.6121 | 94.6565 | 95.0455 | -0.19078 | 0.3509 | -0.3864 | -11.342 | -11.4947 |  | -0.0093772 | 0.55106 | 0.54168 | -0.54168
2038.6122 | 96.3573 | 96.7941 | 96.0543 | 94.7766 | 95.2279 | -0.14149 | 0.4368 | -0.303 | -11.5807 | -11.7077 |  | -0.010536 | 0.58883 | 0.57829 | -0.57829
2098.3859 | 96.6701 | 97.1488 | 96.4064 | 94.891 | 95.4024 | -0.13894 | 0.4787 | -0.2637 | -11.7791 | -11.8853 |  | -0.01184 | 0.62948 | 0.61764 | -0.61764
2159.9122 | 96.5932 | 97.1265 | 96.4209 | 95.1791 | 95.765 | -0.12666 | 0.5333 | -0.1723 | -11.4141 | -11.4882 |  | -0.013306 | 0.67327 | 0.65996 | -0.65996
2223.2424 | 96.279 | 96.9002 | 96.2142 | 95.3218 | 95.866 | -0.084314 | 0.6212 | -0.0648 | -10.9572 | -11.1185 |  | -0.014955 | 0.72047 | 0.70551 | -0.70551
2288.4296 | 96.1507 | 96.812 | 96.125 | 94.9915 | 95.5494 | -0.093282 | 0.6613 | -0.0257 | -11.1592 | -11.3559 |  | -0.01681 | 0.77139 | 0.75458 | -0.75458
2355.5281 | 95.9796 | 96.7292 | 95.8361 | 94.4007 | 95.114 | -0.057876 | 0.7496 | -0.1435 | -11.5789 | -11.6731 |  | -0.018897 | 0.82637 | 0.80748 | -0.80748
2424.594 | 96.2465 | 97.148 | 96.0778 | 94.5778 | 95.4078 | 0.036965 | 0.9015 | -0.1687 | -11.6687 | -11.7032 |  | -0.021246 | 0.88578 | 0.86453 | -0.86453
2495.6849 | 96.336 | 97.2975 | 96.2552 | 94.7101 | 95.6526 | 0.035366 | 0.9615 | -0.0808 | -11.6259 | -11.6095 |  | -0.023889 | 0.95002 | 0.92613 | -0.92613
2568.8603 | 96.6455 | 97.6221 | 96.5505 | 95.2534 | 96.3213 | -0.016083 | 0.9766 | -0.095 | -11.3921 | -11.3169 |  | -0.026863 | 1.0195 | 0.99268 | -0.99268
2644.1812 | 97.0185 | 98.0761 | 96.7368 | 95.3514 | 96.3927 | -0.0070303 | 1.0576 | -0.2817 | -11.6671 | -11.6904 |  | -0.030212 | 1.0948 | 1.0646 | -1.0646
2721.7106 | 96.3693 | 97.5087 | 96.2321 | 94.9241 | 96.17 | -0.0030647 | 1.1394 | -0.1372 | -11.4452 | -11.3418 |  | -0.033981 | 1.1764 | 1.1425 | -1.1425
2801.5132 | 96.0816 | 97.5588 | 95.9468 | 94.8022 | 96.2011 | 0.25048 | 1.4772 | -0.1348 | -11.2794 | -11.1072 |  | -0.038225 | 1.2649 | 1.2267 | -1.2267
2883.6557 | 96.1045 | 97.5791 | 95.8176 | 94.4844 | 95.8966 | 0.15663 | 1.4746 | -0.2869 | -11.6201 | -11.5259 |  | -0.043004 | 1.361 | 1.318 | -1.318
2968.2067 | 96.19 | 97.8346 | 96.1302 | 94.6156 | 96.13 | 0.22777 | 1.6446 | -0.0598 | -11.5744 | -11.4768 |  | -0.048386 | 1.4652 | 1.4168 | -1.4168
3055.2368 | 96.4484 | 98.314 | 96.1894 | 94.6413 | 96.3316 | 0.34162 | 1.8656 | -0.259 | -11.8071 | -11.6408 |  | -0.054447 | 1.5784 | 1.524 | -1.524
3144.8186 | 96.0973 | 98.1753 | 95.7467 | 94.176 | 95.9348 | 0.4379 | 2.078 | -0.3506 | -11.9213 | -11.8026 |  | -0.061273 | 1.7014 | 1.6401 | -1.6401
3237.0271 | 96.2928 | 98.5158 | 96.2376 | 94.2448 | 94.9998 | 0.45708 | 2.223 | -0.0552 | -12.048 | -13.0589 |  | -0.068962 | 1.8349 | 1.7659 | -1.7659
3331.9392 | 96.6524 | 99.0068 | 96.3597 | 94.231 | 96.9471 | 0.45221 | 2.3544 | -0.2927 | -12.4214 | -11.6075 |  | -0.077624 | 1.9798 | 1.9022 | -1.9022
3429.6342 | 96.595 | 98.9535 | 95.1728 | 94.3908 | 96.7969 | 0.30889 | 2.3585 | -1.4222 | -12.2042 | -11.8477 |  | -0.087381 | 2.137 | 2.0496 | -2.0496
3530.1937 | 96.6696 | 99.2605 | 95.5587 | 94.4384 | 96.8257 | 0.38203 | 2.5909 | -1.111 | -12.2312 | -12.0528 |  | -0.098374 | 2.3072 | 2.2089 | -2.2089
3633.7016 | 96.6038 | 99.2448 | 95.9445 | 93.9028 | 96.5151 | 0.26047 | 2.641 | -0.6593 | -12.701 | -12.4692 |  | -0.11076 | 2.4913 | 2.3805 | -2.3805
3740.2445 | 96.3525 | 99.1731 | 95.67 | 93.1914 | 96.2414 | 0.25562 | 2.8206 | -0.6825 | -13.1611 | -12.6761 |  | -0.12471 | 2.6897 | 2.565 | -2.565
3849.9113 | 96.6096 | 99.5055 | 95.7891 | 93.1409 | 96.4476 | 0.13358 | 2.8959 | -0.8205 | -13.4687 | -12.9243 |  | -0.14042 | 2.9027 | 2.7623 | -2.7623
3962.7936 | 95.919 | 98.9153 | 95.1964 | 93.3183 | 96.6153 | 0.024079 | 2.9963 | -0.7226 | -12.6007 | -12.2759 |  | -0.15812 | 3.1303 | 2.9722 | -2.9722
4078.9858 | 95.8896 | 99.3319 | 94.7703 | 93.1045 | 96.7017 | 0.24855 | 3.4423 | -1.1193 | -12.7851 | -12.3817 |  | -0.17805 | 3.3718 | 3.1938 | -3.1938
4198.5847 | 95.6453 | 98.9769 | 95.0076 | 92.361 | 96.3891 | -0.093522 | 3.3316 | -0.6377 | -13.2843 | -12.6813 |  | -0.20049 | 3.6256 | 3.4251 | -3.4251
4321.6904 | 94.8245 | 98.0141 | 94.6241 | 92.138 | 95.8234 | -0.4738 | 3.1896 | -0.2004 | -12.6865 | -12.6645 |  | -0.22574 | 3.8891 | 3.6634 | -3.6634
4448.4057 | 93.4881 | 96.9663 | 93.7756 | 90.9751 | 95.1383 | -0.42601 | 3.4782 | 0.2875 | -12.513 | -12.254 |  | -0.25416 | 4.1584 | 3.9042 | -3.9042
4578.8363 | 92.2915 | 96.1168 | 92.5637 | 89.6063 | 93.6271 | -0.31609 | 3.8253 | 0.2722 | -12.6852 | -12.8058 |  | -0.28612 | 4.4275 | 4.1414 | -4.1414
4713.0912 | 93.174 | 97.6047 | 94.0602 | 90.6653 | 95.2807 | 0.063959 | 4.4307 | 0.8862 | -12.5087 | -12.26 |  | -0.32204 | 4.6888 | 4.3667 | -4.3667
4851.2827 | 93.3308 | 98.0303 | 93.9832 | 90.274 | 95.5014 | 0.12948 | 4.6995 | 0.6524 | -13.0568 | -12.3994 |  | -0.3624 | 4.9324 | 4.57 | -4.57
4993.5259 | 93.6996 | 98.6675 | 94.4165 | 91.5787 | 96.7554 | 0.22866 | 4.9679 | 0.7169 | -12.1209 | -11.6834 |  | -0.40769 | 5.1469 | 4.7392 | -4.7392
5139.9399 | 93.2765 | 98.5955 | 93.8032 | 90.8336 | 96.1763 | 0.45756 | 5.319 | 0.5267 | -12.4429 | -11.9616 |  | -0.45848 | 5.3199 | 4.8614 | -4.8614
5290.6469 | 93.8268 | 98.8472 | 94.0799 | 91.407 | 96.8915 | 0.096231 | 5.0204 | 0.2531 | -12.4198 | -11.8595 |  | -0.51538 | 5.4395 | 4.9242 | -4.9242
5445.7726 | 93.8879 | 98.7861 | 94.5246 | 91.6824 | 96.89 | -0.018979 | 4.8982 | 0.6367 | -12.2055 | -11.9151 |  | -0.57903 | 5.4962 | 4.9172 | -4.9172
5605.4468 | 94.1351 | 99.2399 | 94.7341 | 91.1228 | 96.0455 | 0.27056 | 5.1048 | 0.599 | -13.0123 | -12.9238 |  | -0.65012 | 5.4844 | 4.8342 | -4.8342
5769.8028 | 94.8104 | 99.2212 | 95.3172 | 90.3404 | 95.1544 | -0.2635 | 4.4108 | 0.5068 | -14.47 | -14.3303 |  | -0.72941 | 5.4037 | 4.6743 | -4.6743
5938.9778 | 95.3401 | 99.7735 | 96.5486 | 91.967 | 96.3222 | -0.0081732 | 4.4334 | 1.2085 | -13.3731 | -13.4595 |  | -0.81764 | 5.2592 | 4.4416 | -4.4416
6113.1131 | 96.0387 | 99.9119 | 96.9326 | 92.0152 | 96.2457 | -0.27143 | 3.8732 | 0.8939 | -14.0235 | -13.9376 |  | -0.91561 | 5.0602 | 4.1446 | -4.1446
6292.3542 | 95.4644 | 98.5919 | 96.6778 | 92.0814 | 96.5748 | -0.66714 | 3.1275 | 1.2134 | -13.383 | -12.6842 |  | -1.0241 | 4.8188 | 3.7946 | -3.7946
6476.8508 | 94.4649 | 97.5593 | 95.8177 | 91.4016 | 95.3497 | -0.3092 | 3.0944 | 1.3528 | -13.0633 | -12.5188 |  | -1.144 | 4.5476 | 3.4036 | -3.4036
6666.757 | 94.0466 | 96.4814 | 94.4763 | 91.3062 | 95.0106 | -0.54803 | 2.4348 | 0.4297 | -12.7404 | -12.0188 |  | -1.2759 | 4.2587 | 2.9828 | -2.9828
6862.2314 | 93.6073 | 95.851 | 94.0982 | 91.9847 | 94.6102 | -0.29832 | 2.2437 | 0.4909 | -11.6226 | -11.5391 |  | -1.4206 | 3.9626 | 2.542 | -2.542
7063.4372 | 94.4574 | 96.4091 | 94.0767 | 91.6208 | 93.5538 | -0.13721 | 1.9517 | -0.3807 | -12.8366 | -12.9925 |  | -1.5786 | 3.6676 | 2.0889 | -2.0889
7270.5426 | 94.6849 | 96.576 | 96.2355 | 92.1276 | 93.5015 | 0.26186 | 1.8911 | 1.5506 | -12.5573 | -12.8126 |  | -1.7506 | 3.3798 | 1.6292 | -1.6292
7483.7204 | 95.048 | 96.7475 | 96.9128 | 91.4932 | 92.3952 | 0.53245 | 1.6995 | 1.8648 | -13.5548 | -13.8198 |  | -1.9367 | 3.1037 | 1.167 | -1.167
7703.1488 | 94.9574 | 95.5898 | 97.4867 | 90.8724 | 91.7122 | -0.072639 | 0.6324 | 2.5293 | -14.085 | -13.9502 |  | -2.137 | 2.8421 | 0.70504 | -0.70504
7929.0109 | 94.6776 | 94.9737 | 96.5926 | 91.1228 | 91.0633 | 0.051165 | 0.2961 | 1.915 | -13.5548 | -13.8592 |  | -2.3516 | 2.5965 | 0.24494 | -0.24494
8161.4956 | 94.24 | 94.6018 | 97.6809 | 89.8247 | 89.8974 | 0.57396 | 0.3618 | 3.4409 | -14.4153 | -14.1304 |  | -2.5798 | 2.3676 | -0.21216 | 0.21216
8400.7968 | 95.4548 | 95.5521 | 97.5549 | 89.8016 | 89.5251 | 0.76278 | 0.0973 | 2.1001 | -15.6532 | -15.2642 |  | -2.821 | 2.1555 | -0.66548 | 0.66548
8647.1146 | 94.9442 | 94.1389 | 97.7007 | 89.7175 | 90.1204 | 0.30906 | -0.8053 | 2.7565 | -15.2267 | -13.7094 |  | -3.074 | 1.9597 | -1.1144 | 1.1144
8900.6545 | 95.173 | 93.4969 | 96.9407 | 89.7479 | 89.0762 | -0.11803 | -1.6761 | 1.7677 | -15.4251 | -14.5387 |  | -3.3376 | 1.7795 | -1.5581 | 1.5581
9161.6285 | 95.1423 | 93.0115 | 96.0882 | 90.4664 | 88.9922 | -0.13509 | -2.1308 | 0.9459 | -14.6759 | -14.1544 |  | -3.6098 | 1.6141 | -1.9957 | 1.9957
9430.2544 | 95.2211 | 92.6221 | 97.0017 | 89.4908 | 87.825 | -0.17288 | -2.599 | 1.7806 | -15.7303 | -14.97 |  | -3.8886 | 1.4625 | -2.4261 | 2.4261
9706.7566 | 95.0672 | 92.6177 | 98.7052 | 88.7066 | 87.7532 | 0.39837 | -2.4495 | 3.638 | -16.3606 | -14.4661 |  | -4.1715 | 1.3236 | -2.8479 | 2.8479
9991.3661 | 95.31 | 92.1019 | 98.4904 | 89.1994 | 87.5017 | 0.051158 | -3.2081 | 3.1804 | -16.1106 | -14.549 |  | -4.4558 | 1.1966 | -3.2593 | 3.2593
10284.3205 | 94.4899 | 90.4693 | 96.9649 | 89.8306 | 87.7902 | -0.36222 | -4.0206 | 2.475 | -14.6593 | -13.0413 |  | -4.7387 | 1.0803 | -3.6584 | 3.6584
10585.8646 | 93.4694 | 89.7944 | 96.6654 | 88.8256 | 86.3949 | 0.36816 | -3.675 | 3.196 | -14.6438 | -13.0313 |  | -5.0172 | 0.97401 | -4.0432 | 4.0432
10896.2502 | 94.0563 | 89.9563 | 95.954 | 88.6123 | 85.1668 | 0.31149 | -4.1 | 1.8977 | -15.444 | -14.478 |  | -5.2883 | 0.87677 | -4.4115 | 4.4115
11215.7365 | 93.5152 | 88.7866 | 95.0463 | 88.1482 | 84.2046 | 0.032691 | -4.7286 | 1.5311 | -15.367 | -14.5493 |  | -5.5491 | 0.78782 | -4.7613 | 4.7613
11544.5904 | 93.1021 | 87.6988 | 95.6711 | 86.2787 | 82.3994 | -0.31265 | -5.4033 | 2.569 | -16.8234 | -15.6121 |  | -5.7971 | 0.70644 | -5.0906 | 5.0906
11883.0866 | 92.7053 | 87.9292 | 97.1424 | 85.7244 | 81.8243 | 0.62181 | -4.7761 | 4.4371 | -16.9809 | -15.4831 |  | -6.0299 | 0.63197 | -5.3979 | 5.3979
12231.5077 | 93.4643 | 87.9868 | 96.5969 | 87.0704 | 83.2695 | 0.20428 | -5.4775 | 3.1326 | -16.3939 | -14.513 |  | -6.2456 | 0.56381 | -5.6818 | 5.6818
12590.1448 | 93.5395 | 87.5684 | 94.9595 | 87.423 | 83.1917 | -0.029706 | -5.9711 | 1.42 | -16.1165 | -14.4064 |  | -6.4428 | 0.50144 | -5.9414 | 5.9414
12959.2974 | 93.3794 | 87.4789 | 95.7577 | 86.4818 | 82.0062 | 0.27585 | -5.9005 | 2.3783 | -16.8976 | -15.1968 |  | -6.6207 | 0.44435 | -6.1764 | 6.1764
13339.2738 | 93.5072 | 86.6317 | 95.8223 | 85.0315 | 80.7819 | -0.48876 | -6.8755 | 2.3151 | -18.4757 | -16.3386 |  | -6.7789 | 0.39211 | -6.3867 | 6.3867
13730.3915 | 92.3516 | 85.7629 | 96.0704 | 86.8553 | 82.117 | -0.015607 | -6.5887 | 3.7188 | -15.4963 | -13.6615 |  | -6.9174 | 0.34433 | -6.5731 | 6.5731
14132.977 | 93.3469 | 87.2289 | 95.3204 | 85.8375 | 81.142 | 0.61836 | -6.118 | 1.9735 | -17.5094 | -15.4685 |  | -7.037 | 0.30066 | -6.7364 | 6.7364
14547.3666 | 94.53 | 87.1435 | 96.0461 | 86.5307 | 81.2488 | -0.50867 | -7.3865 | 1.5161 | -17.9993 | -16.4034 |  | -7.1386 | 0.26079 | -6.8778 | 6.8778
14973.9064 | 93.7846 | 86.5771 | 96.3408 | 84.6197 | 81.421 | -0.20843 | -7.2075 | 2.5562 | -19.1649 | -15.3645 |  | -7.2235 | 0.22444 | -6.9991 | 6.9991
15412.9527 | 94.1679 | 87.2461 | 96.5808 | 85.4477 | 81.7084 | 0.18 | -6.9218 | 2.4129 | -18.7202 | -15.3577 |  | -7.2932 | 0.19137 | -7.1018 | 7.1018
15864.8722 | 91.7537 | 84.8896 | 94.5315 | 84.4975 | 79.6576 | 0.32377 | -6.8641 | 2.7778 | -17.2562 | -14.9082 |  | -7.3492 | 0.16136 | -7.1879 | 7.1879
16330.0423 | 90.9735 | 84.9355 | 93.4721 | 82.574 | 77.6469 | 1.2211 | -6.038 | 2.4986 | -18.3995 | -16.0675 |  | -7.3934 | 0.13423 | -7.2591 | 7.2591
16808.8515 | 91.3878 | 84.4678 | 93.0757 | 81.5143 | 76.7897 | 0.39741 | -6.92 | 1.6879 | -19.8735 | -17.2807 |  | -7.4272 | 0.10983 | -7.3174 | 7.3174
17301.6998 | 90.3575 | 82.2432 | 91.7275 | 80.3568 | 76.7691 | -0.74983 | -8.1143 | 1.37 | -20.0007 | -16.2239 |  | -7.4525 | 0.088017 | -7.3645 | 7.3645
17808.9988 | 88.8543 | 82.2367 | 92.1451 | 80.4666 | 76.9016 | 0.78434 | -6.6176 | 3.2908 | -18.3877 | -14.5508 |  | -7.4706 | 0.068694 | -7.4019 | 7.4019
18331.1722 | 89.9654 | 83.4937 | 92.28 | 79.7281 | 76.0281 | 0.95962 | -6.4717 | 2.3146 | -20.2373 | -16.506 |  | -7.4831 | 0.051781 | -7.4313 | 7.4313
18868.656 | 88.8966 | 81.4609 | 90.1407 | 79.558 | 75.4147 | 0.018236 | -7.4357 | 1.2441 | -19.3386 | -16.028 |  | -7.4912 | 0.037227 | -7.4539 | 7.4539
19421.8994 | 87.4177 | 81.068 | 89.9196 | 78.3855 | 74.0158 | 1.1213 | -6.3497 | 2.5019 | -19.0322 | -15.9309 |  | -7.496 | 0.025008 | -7.471 | 7.471
19991.3642 | 88.7144 | 82.6137 | 90.1202 | 78.0833 | 74.2577 | 1.3827 | -6.1007 | 1.4058 | -20.6311 | -16.9733 |  | -7.4985 | 0.015132 | -7.4834 | 7.4834
20577.5262 | 89.449 | 82.8578 | 90.1056 | 75.7392 | 74.534 | 0.90066 | -6.5912 | 0.6566 | -23.7098 | -17.4231 |  | -7.4996 | 0.0077369 | -7.4919 | 7.4919
21180.8749 | 88.399 | 81.9052 | 88.3863 | 79.2338 | 76.9721 | 1.0035 | -6.4938 | -0.0127 | -19.1652 | -13.9296 |  | -7.4999 | 0.0026787 | -7.4973 | 7.4973
21801.9142 | 87.8054 | 81.7301 | 89.6893 | 76.2404 | 74.5181 | 1.4245 | -6.0753 | 1.8839 | -21.565 | -15.7875 |  | -7.5 | 0.00021759 | -7.4998 | 7.4998
22441.1629 | 88.0838 | 80.1418 | 88.7565 | 74.0187 | 72.4748 | -0.44254 | -7.942 | 0.6727 | -24.0651 | -18.1095 |  | -7.5 | 0.00054117 | -7.4995 | 7.4995
23099.1548 | 88.7755 | 82.5908 | 90.8484 | 76.7172 | 73.0496 | 1.3113 | -6.1847 | 2.0729 | -22.0583 | -18.2299 |  | -7.4999 | 0.0039094 | -7.496 | 7.496
23776.4396 | 90.5348 | 83.1042 | 92.0156 | 75.3169 | 72.2896 | 0.057972 | -7.4306 | 1.4808 | -25.2179 | -20.7566 |  | -7.4992 | 0.010674 | -7.4886 | 7.4886
24473.5828 | 91.1586 | 82.4595 | 91.3241 | 77.2925 | 74.2162 | -1.2234 | -8.6991 | 0.1655 | -23.8661 | -19.4667 |  | -7.4971 | 0.021308 | -7.4757 | 7.4757
25191.1669 | 91.6351 | 83.1571 | 92.8445 | 77.3918 | 71.5137 | -1.0229 | -8.478 | 1.2094 | -24.2433 | -22.6663 |  | -7.4915 | 0.036444 | -7.4551 | 7.4551
25929.791 | 93.1738 | 82.0179 | 94.0522 | 76.1452 | 71.5397 |  | -11.1559 | 0.8784 | -27.0286
26690.0722 | 94.4673 | 82.8507 | 95.0523 | 77.6108 | 72.9303 |  | -11.6166 | 0.585 | -26.8565
27472.6453 | 96.8139 | 83.7102 | 97.0748 | 79.7434 | 72.1527 |  | -13.1037 | 0.2609 | -27.0705
28278.1642 | 98.8835 | 83.7753 | 99.9798 | 81.7811 | 72.6164 |  | -15.1082 | 1.0963 | -27.1024
29107.3014 | 102.6924 | 85.7222 | 103.2887 | 84.0636 | 73.5469 |  | -16.9702 | 0.5963 | -28.6288
29960.7496 | 106.2124 | 87.0301 | 106.1544 | 87.2178 | 75.3428 |  | -19.1823 | -0.058 | -28.9946
30839.2215 | 103.1169 | 82.5805 | 103.0418 | 82.3617 | 70.1046 |  | -20.5364 | -0.0751 | -30.7552
31743.4509 | 98.0325 | 75.4841 | 97.7387 | 76.1538 | 62.4016 |  | -22.5484 | -0.2938 | -31.8787
32674.193 | 93.705 | 70.439 | 92.7925 | 70.412 | 56.5221 |  | -23.266 | -0.9125 | -33.293
33632.2251 | 89.8094 | 66.348 | 89.5571 | 64.1657 | 50.7335 |  | -23.4614 | -0.2523 | -35.6437
34618.3475 | 87.0684 | 64.0744 | 86.7744 | 61.1634 | 50.2995 |  | -22.994 | -0.294 | -35.905
35633.3837 | 85.8879 | 64.0755 | 85.1713 | 58.091 | 49.4522 |  | -21.8124 | -0.7166 | -37.7969
36678.1816 | 84.8934 | 63.6601 | 84.2528 | 57.262 | 49.9022 |  | -21.2333 | -0.6406 | -37.6314
37753.6137 | 82.6604 | 63.5142 | 81.7817 | 53.9843 | 50.5047 |  | -19.1462 | -0.8787 | -38.6761
38860.5783 | 80.4422 | 62.9019 | 78.2089 | 51.2553 | 50.6353 |  | -17.5403 | -2.2333 | -39.1869

## distortion
Hz | OpenRefMic | EMM-6 | OpenRefMicestimated measurement noise floor | EMM-6 estimated measurement noise floor
20 | -24.6939 | -21.2643 | -22.5128 | -21.4959
20.5865 | -24.4144 | -20.641 | -23.2765 | -21.8902
21.1903 | -23.7487 | -19.9623 | -24.1884 | -22.3883
21.8117 | -22.7435 | -19.2332 | -25.3085 | -23.0178
22.4514 | -21.5124 | -18.4903 | -26.7491 | -23.9054
23.1098 | -20.4397 | -17.8453 | -28.4329 | -24.972
23.7876 | -19.5238 | -17.2981 | -30.3662 | -26.2628
24.4852 | -18.8044 | -16.8726 | -32.3282 | -27.839
25.2033 | -18.3354 | -16.6154 | -33.2443 | -29.6221
25.9424 | -18.0757 | -16.5108 | -33.0434 | -31.3741
26.7032 | -18.0266 | -16.5698 | -32.1961 | -32.7024
27.4863 | -18.1977 | -16.8082 | -31.3449 | -33.1995
28.2924 | -18.6025 | -17.2438 | -30.8997 | -32.9223
29.1222 | -19.2546 | -17.8903 | -31.0111 | -32.4647
29.9762 | -20.1585 | -18.7484 | -31.7484 | -32.2982
30.8553 | -21.2621 | -19.7634 | -33.0894 | -32.5695
31.7602 | -22.416 | -20.8002 | -34.8745 | -33.2368
32.6917 | -23.2886 | -21.5997 | -36.52 | -34.156
33.6504 | -23.4247 | -21.8317 | -36.9284 | -35.0396
34.6373 | -22.6804 | -21.3656 | -36.023 | -35.5215
35.6531 | -21.4813 | -20.479 | -34.9612 | -35.4957
36.6987 | -20.2288 | -19.4859 | -34.3303 | -35.1124
37.7749 | -19.1216 | -18.5758 | -34.2437 | -34.5817
38.8828 | -18.2261 | -17.8283 | -34.6329 | -34.0915
40.0231 | -17.56 | -17.2734 | -35.3012 | -33.8448
41.1968 | -17.163 | -16.9534 | -35.8637 | -34.1172
42.405 | -17.0042 | -16.849 | -36.1723 | -34.8871
43.6486 | -17.1034 | -16.9817 | -36.4806 | -35.7717
44.9287 | -17.5403 | -17.4282 | -37.2445 | -35.6055
46.2463 | -18.291 | -18.1649 | -38.2255 | -34.6287
47.6026 | -19.3789 | -19.2059 | -38.3596 | -34.0902
48.9986 | -20.7054 | -20.4454 | -37.8385 | -34.7118
50.4356 | -22.0392 | -21.6519 | -38.2578 | -36.7694
51.9147 | -23.1524 | -22.6316 | -40.2676 | -39.3805
53.4372 | -23.8997 | -23.311 | -42.6664 | -40.3888
55.0044 | -24.4332 | -23.8945 | -41.9071 | -40.0314
56.6175 | -24.944 | -24.5758 | -40.7044 | -39.4012
58.2779 | -25.5654 | -25.4591 | -41.3069 | -38.5873
59.987 | -26.359 | -26.5329 | -43.7555 | -38.1914
61.7463 | -27.3443 | -27.6952 | -45.9976 | -39.1174
63.5571 | -28.5409 | -28.888 | -45.6732 | -41.8528
65.421 | -29.9927 | -30.1764 | -45.3843 | -45.7351
67.3396 | -31.679 | -31.5705 | -46.1891 | -45.4634
69.3145 | -33.4024 | -32.8389 | -46.9689 | -43.0751
71.3473 | -34.9424 | -33.9189 | -47.4255 | -42.1749
73.4397 | -36.3392 | -35.2894 | -48.0367 | -42.9835
75.5934 | -37.7355 | -37.3624 | -48.2719 | -46.0373
77.8104 | -39.1708 | -39.6332 | -48.5586 | -50.9979
80.0923 | -40.8427 | -41.2822 | -49.3903 | -50.155
82.4412 | -42.9188 | -42.5246 | -48.8273 | -49.2649
84.8589 | -44.2792 | -43.4424 | -46.1846 | -52.2925
87.3476 | -43.3245 | -44.5532 | -44.0711 | -52.8067
89.9092 | -41.496 | -46.0626 | -44.3818 | -50.155
92.546 | -39.7126 | -43.9181 | -47.7515 | -48.5747
95.2601 | -39.368 | -41.3343 | -51.9569 | -47.2369
98.0537 | -42.2386 | -41.7099 | -51.461 | -46.9476
100.9294 | -43.2511 | -42.7085 | -48.5802 | -47.6884
103.8893 | -42.0811 | -41.7338 | -47.815 | -50.797
106.9361 | -46.9214 | -44.0401 | -50.4082 | -52.0883
110.0722 | -49.1137 | -46.1801 | -53.1305 | -48.6297
113.3003 | -49.7006 | -45.5625 | -51.5177 | -49.2062
116.623 | -46.9115 | -46.5319 | -51.565 | -51.8812
120.0432 | -45.7321 | -46.1378 | -54.1553 | -54.1903
123.5637 | -46.3094 | -44.4949 | -54.8517 | -54.5991
127.1875 | -50.5943 | -47.9365 | -52.7388 | -52.9463
130.9175 | -51.7673 | -50.8191 | -49.1545 | -53.0511
134.7569 | -51.2677 | -49.997 | -51.2114 | -55.7917
138.7089 | -53.8447 | -53.351 | -53.6523 | -51.2109
142.7768 | -51.751 | -49.9055 | -51.8951 | -49.1679
146.964 | -51.6099 | -51.6533 | -47.3159 | -49.8764
151.274 | -54.1679 | -48.1325 | -45.8705 | -50.1785
155.7105 | -50.4832 | -48.2645 | -51.4371 | -51.3545
160.277 | -52.8985 | -49.983 | -45.7863 | -50.655
164.9774 | -52.0174 | -52.5001 | -49.4848 | -51.6988
169.8157 | -50.3986 | -54.0876 | -54.7765 | -52.3967
174.7959 | -52.5356 | -50.7022 | -54.9897 | -50.1149
179.9221 | -49.5424 | -51.0035 | -56.9822 | -51.0901
185.1987 | -54.7869 | -51.6835 | -55.525 | -50.6582
190.63 | -51.4712 | -49.5584 | -55.8721 | -52.2113
196.2206 | -50.9795 | -50.3612 | -53.9743 | -53.154
201.9751 | -52.0167 | -51.1986 | -54.6499 | -47.852
207.8985 | -53.5109 | -51.2005 | -53.9895 | -52.3522
213.9955 | -52.4239 | -51.7373 | -54.8131 | -49.2692
220.2713 | -52.0844 | -49.9306 | -50.1132 | -51.9432
226.7312 | -50.8254 | -51.0129 | -54.8468 | -48.7769
233.3805 | -53.1096 | -50.5429 | -51.442 | -50.5569
240.2249 | -54.3458 | -50.3295 | -54.614 | -50.1312
247.27 | -52.6659 | -47.9894 | -48.1292 | -47.5174
254.5216 | -56.7964 | -50.6871 | -50.3985 | -49.1138
261.986 | -49.2354 | -47.628 | -50.474 | -50.1351
269.6692 | -50.3667 | -50.9114 | -52.6991 | -49.4646
277.5778 | -55.6102 | -50.7848 | -52.077 | -53.7596
285.7183 | -54.7389 | -50.7339 | -52.8714 | -51.1113
294.0976 | -52.5963 | -51.9915 | -54.6514 | -51.1341
302.7226 | -49.7847 | -49.9607 | -55.2811 | -50.2541
311.6005 | -51.0593 | -50.6065 | -54.003 | -52.44
320.7388 | -48.8328 | -50.6532 | -49.14 | -48.7725
330.1451 | -52.476 | -51.7929 | -55.4793 | -50.9571
339.8272 | -49.2681 | -47.9052 | -52.4363 | -49.2568
349.7933 | -52.1909 | -47.9696 | -50.3268 | -49.2262
360.0517 | -50.8037 | -50.1839 | -50.2441 | -49.309
370.6109 | -52.7985 | -49.8622 | -60.1244 | -49.4602
381.4798 | -52.9239 | -49.7808 | -55.1091 | -52.941
392.6675 | -53.1942 | -48.4362 | -53.5113 | -48.7424
404.1832 | -51.374 | -50.5524 | -57.7534 | -51.786
416.0367 | -53.4712 | -51.5924 | -59.6126 | -50.1417
428.2378 | -52.474 | -51.2689 | -60.0968 | -51.529
440.7967 | -56.1447 | -51.5324 | -54.3657 | -52.1246
453.7239 | -55.4613 | -50.6726 | -52.8502 | -54.3955
467.0303 | -56.867 | -51.8938 | -60.6669 | -52.2274
480.7268 | -57.2309 | -51.5025 | -59.1544 | -49.4993
494.8251 | -55.3278 | -47.6975 | -54.9206 | -49.349
509.3368 | -58.0081 | -46.3556 | -55.6081 | -50.6696
524.2741 | -57.6919 | -53.1211 | -63.8253 | -52.4634
539.6495 | -60.409 | -50.2313 | -58.1505 | -51.2044
555.4758 | -58.11 | -51.9236 | -60.9114 | -53.1427
571.7662 | -57.3586 | -50.5831 | -54.951 | -51.2881
588.5343 | -57.9797 | -51.7203 | -62.7555 | -52.3545
605.7943 | -54.91 | -52.0749 | -63.7248 | -53.4185
623.5604 | -59.4692 | -50.9872 | -64.9217 | -51.9384
641.8475 | -55.0404 | -51.274 | -68.1163 | -52.2283
660.6709 | -59.7305 | -52.9011 | -63.1988 | -50.3533
680.0464 | -57.6735 | -50.2676 | -60.4447 | -51.7127
699.9901 | -55.3619 | -51.7842 | -64.7506 | -51.8484
720.5187 | -57.216 | -51.6723 | -64.9657 | -51.8988
741.6493 | -55.9544 | -52.2008 | -61.1118 | -49.8029
763.3996 | -55.2772 | -47.2292 | -63.7095 | -51.7841
785.7878 | -55.5657 | -48.1425 | -60.5976 | -50.4472
808.8326 | -58.7826 | -49.5843 | -61.1584 | -51.5006
832.5532 | -55.4965 | -50.0215 | -62.0164 | -51.3378
856.9695 | -56.7891 | -50.3744 | -62.3034 | -50.5749
882.1018 | -57.6391 | -48.5974 | -62.1494 | -50.7615
907.9711 | -54.6722 | -51.1906 | -64.6127 | -48.3433
934.5992 | -56.4607 | -47.536 | -61.766 | -48.5323
962.0081 | -56.8298 | -48.7379 | -60.8141 | -47.4948
990.2209 | -56.9954 | -47.3882 | -59.6487 | -48.0873
1019.2611 | -57.3499 | -49.2904 | -59.3804 | -50.1646
1049.1529 | -58.0252 | -48.7896 | -62.0184 | -51.2351
1079.9214 | -55.1433 | -48.1583 | -59.0554 | -47.3371
1111.5922 | -57.7294 | -49.8621 | -65.1734 | -51.4696
1144.1918 | -61.3763 | -51.0193 | -64.5342 | -49.6108
1177.7475 | -56.5956 | -50.7919 | -66.6325 | -49.1298
1212.2872 | -57.1429 | -51.7811 | -63.5753 | -49.8945
1247.8399 | -58.0844 | -52.516 | -66.8158 | -54.1222
1284.4353 | -54.2246 | -49.1437 | -69.1861 | -50.5976
1322.1039 | -56.8231 | -50.9128 | -63.2192 | -52.8859
1360.8772 | -54.7206 | -49.628 | -64.0318 | -52.9541
1400.7876 | -57.9488 | -50.2982 | -63.8278 | -50.2154
1441.8684 | -57.2172 | -47.8456 | -61.2522 | -48.4072
1484.154 | -58.2013 | -52.2444 | -65.9561 | -50.3427
1527.6798 | -56.5184 | -47.8292 | -64.058 | -49.8726
1572.482 | -55.9588 | -49.4634 | -68.0235 | -48.7188
1618.5981 | -54.3298 | -47.8035 | -62.5223 | -49.5597
1666.0666 | -57.3374 | -49.9704 | -63.286 | -50.727
1714.9273 | -59.6114 | -48.6044 | -62.1553 | -52.8755
1765.2209 | -56.9402 | -48.0468 | -64.5008 | -49.0587
1816.9895 | -59.6477 | -49.316 | -65.1152 | -50.1188
1870.2763 | -59.5582 | -48.1292 | -69.9971 | -51.5272
1925.1258 | -57.105 | -49.407 | -67.0492 | -52.6799
1981.5839 | -61.5436 | -48.6493 | -63.8678 | -50.0739
2039.6977 | -61.1803 | -49.9221 | -62.0011 | -52.3643
2099.5159 | -56.1949 | -50.9341 | -63.768 | -51.3979
2161.0883 | -63.3783 | -49.2041 | -66.7558 | -50.3404
2224.4664 | -61.6436 | -49.5051 | -64.5616 | -48.6878
2289.7033 | -57.6089 | -51.5118 | -71.8889 | -49.952
2356.8533 | -57.483 | -50.2794 | -64.859 | -48.6707
2425.9727 | -55.9025 | -47.7173 | -66.0846 | -47.1871
2497.1191 | -54.3127 | -51.3877 | -67.4678 | -50.6881
2570.352 | -57.8694 | -49.2851 | -65.8568 | -48.8032
2645.7326 | -58.005 | -50.211 | -70.656 | -51.8975
2723.3239 | -57.714 | -50.1826 | -70.3457 | -53.3868
2803.1907 | -58.7401 | -50.136 | -67.7675 | -49.3955
2885.3998 | -59.1 | -50.7606 | -66.3513 | -48.7076
2970.0198 | -54.8571 | -49.0078 | -66.3008 | -48.1373
3057.1215 | -56.5726 | -48.7912 | -66.9134 | -51.9055
3146.7776 | -57.6285 | -51.1208 | -73.9232 | -51.2393
3239.063 | -55.8519 | -54.416 | -70.8359 | -55.1071
3334.0549 | -56.6186 | -49.3002 | -78.1201 | -50.8293
3431.8326 | -59.9507 | -51.016 | -73.5009 | -54.8137
3532.4778 | -56.7279 | -49.6403 | -74.0081 | -50.6576
3636.0746 | -55.8554 | -47.9127 | -70.7181 | -49.8467
3742.7097 | -56.5454 | -47.9952 | -68.6514 | -52.5843
3852.472 | -58.2775 | -51.6533 | -69.5197 | -46.6561
3965.4533 | -60.8859 | -47.3099 | -67.5019 | -51.0326
4081.748 | -60.3556 | -46.2427 | -72.6002 | -51.7939
4201.4532 | -57.8264 | -48.5237 | -72.2634 | -50.1285
4324.6691 | -55.0937 | -54.9647 | -72.7843 | -50.5639
4451.4985 | -56.3642 | -50.9611 | -69.9365 | -51.75
4582.0474 | -55.9172 | -50.8686 | -75.8075 | -52.792
4716.425 | -55.2057 | -51.5837 | -74.0273 | -48.9731
4854.7434 | -53.8041 | -48.6308 | -79.3701 | -54.5941
4997.1183 | -53.5187 | -49.5146 | -74.5424 | -54.5102
5143.6686 | -53.285 | -52.4773 | -76.3362 | -55.5538
5294.5167 | -53.2308 | -47.4299 | -75.083 | -48.8802
5449.7888 | -55.5441 | -46.8434 | -78.0012 | -50.4549
5609.6146 | -55.6316 | -49.0699 | -65.5541 | -46.114
5774.1275 | -58.7534 | -45.7095 | -63.0645 | -43.5081
5943.4652 | -51.7107 | -37.5462 | -64.1641 | -41.3563
6117.7689 | -53.7323 | -42.803 | -71.2284 | -49.943
6297.1845 | -56.072 | -45.6564 | -70.4086 | -47.7244
6481.8618 | -56.5727 | -47.2979 | -73.009 | -52.6002
6671.9551 | -58.7229 | -47.3207 | -70.8091 | -52.4036
6867.6233 | -57.7838 | -57.8109 | -71.0939 | -54.1861
7069.0298 | -58.5103 | -47.227 | -66.6881 | -50.7631
7276.343 | -45.0428 | -32.1984 | -50.2311 | -33.7995
7489.736 | -49.8533 | -41.6719 | -55.6889 | -41.2148
7709.3872 | -53.9624 | -51.9064 | -59.1573 | -47.7058
7935.4801 | -62.6003 | -45.6955 | -69.1636 | -48.9307
8168.2037 | -67.9331 | -50.0247 | -61.9074 | -55.2456
8407.7523 | -64.3451 | -47.6054 | -69.8936 | -44.2977
8654.3261 | -45.5053 | -35.7968 | -65.7952 | -39.4345
8908.1312 | -48.2762 | -36.5316 | -46.3979 | -40.3245
9169.3796 | -67.4324 | -42.9244 | -55.7043 | -50.869
9438.2897 | -52.5069 | -47.291 | -57.3514 | -46.6247
9715.086 | -67.3391 | -47.2357 | -69.2699 | -47.8291
10000 | -58.0028 | -55.5691 | -66.532 | -57.4572